GENERAL

  - Section includes

    - Amperometric chlorine analyzer for continuous measurement of free or total chlorine in aqueous solutions.

  - Measurement Procedures

    - The method of measuring free or total chlorine will be with a three-electrode amperometric sensor immersed into an electrolytic medium with a membrane, selective to chlorine, separating it from the sample.

  - Alternates

    - Other methods of chlorine measurement, such as two-electrode amperometric, open cell amperometric, or measurements with external pH compensation are not acceptable.

  - System Description

    - Performance Requirements

      - Measurement range: 0 to 20 ppm chlorine for either free or total chlorine

      - Total Chlorine

        - Low Limit Of Detection (LOD):  30 ppb (0.03 ppm) or better

        - Limit Of Quantitation (LOQ):  90 ppb (0.09 ppm) or better

        - Repeatability/precision:  30 ppb or 3%, whichever is greater

        - Response time: ~100 s for 90% change (T90) (At a stable T and pH)

        - Interference: Chlorine Dioxide, Ozone, and chalk deposits

      - Free Chlorine

        - Low Limit Of Detection (LOD):  30 ppb (0.03 ppm) or better

        - Limit Of Quantitation (LOQ):  90 ppb (0.09 ppm) or better

        - Repeatability/precision:  30 ppb or 3%, whichever is greater

        - Response time: ~140 s for 90% change (T90) (At a stable T and pH)

        - Interference:  Monochloramine, Chlorine Dioxide, Ozone, and chalk deposits

      - Drift: <10% with regular calibration (calibration will be weekly to quarterly depending on the application, given stable sample temperature and pH of water sample)

      - Specificity/Selectivity:  Non-specific to a certain chlorine form, responds to any chlorine species and other oxidizers as noted in the interference section

      - Calibration method:  Customer has the option to use one (zero or slope) or two point (zero and slope) calibration. Two point calibration with chemical zero is recommended for chlorine concentration <0.5 ppm

      - Verification procedure:  One-point process calibration (slope) against a standard reference method.

  - Certifications (when connected to a Hach sc controller):

    - EMC: CE compliant for conducted and radiated emissions CISPR 11 (Class A limits), EMC Immunity EN 61326-1 (Industrial limits)

    - Safety: General Purpose UL/CSA 61010-1 with cETLus safety mark

    - Australian C-TICK and Korean KC Markings

    - NEMA 4X/IP65 Water and Dust Ingress Ratings

  - Environmental Requirements

    - Operational Criteria

      - Operating temperature: 5 to 45 °C (41 to 113 °F)

      - Relative humidity: 0-95%, non condensing

    - Sample Requirements

      - Maximum back pressure the chlorine sensor can manage without failure:

        - 0.5 bar, no pressure impulses and/or vibrations

      - Temperature:  0 to 45 °C (33 to 113 °F)

      - Temperature compensation range:  5 to 45 °C (41 to 113 °F)

      - Flow:  30-50 L/hr, 40 L/hr - optimal (7.9-13.2 g/hr, 10.6 g/hr - optimal)

    - Storage Requirements

      - Electrolyte: 15 to 25°C (59 to 77°F)

      - Chlorine sensors: 0 to 50°C (32 to 122°F) dry without electrolyte

      - Panel: -20 to 60°C (-4 to 149°F)

  - Warranty

    - The product includes a one-year warranty from the date of shipment.

  - Maintenance Service

    - Scheduled maintenance:

      - Calibration by comparison with lab method: every 1 month or as necessary

      - Replace membrane: every 6 months or as necessary

      - Replace electrolyte: every 3 to 6 months or as necessary

    - Unscheduled maintenance

      - Cleaning as needed based on environmental conditions.

      - Working electrode tip polishing when required

PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Model CLF10sc Reagentless Free Chlorine Analyzer

      - Model CLT10sc Reagentless Total Chlorine Analyzer

  - Manufactured Unit

    - The CLF10 sc or CLT10sc analyzer consists of:

      - Three-Electrode Amperometric Chlorine sensor

      - Chlorine sensor flow cell with integrated flow sensor

      - pH flow cell with grab sample port

      - Digital gateway for communication between probes and controller

      - Stainless steel panel

  - Equipment

    - The CLF10 sc and CLT10 sc work with Hach sc model controllers only. (Specific controller specifications can be found in the associated sensor CSI specifications)

    - The amperometric cell of the sensor consists of:

      - Gold cathode

      - Stainless steel counter electrode

      - Silver/silver chloride reference electrode

      - pH buffered electrolyte

      - Sensor membrane to filter chlorine species selectively and to provide interface between the electrochemical cell and the sample

    - Wetted materials as follows:

      - Chlorine Measuring Cell: PVC

      - Chlorine Sensor Body: PVC

      - Chlorine Sensor Flow Cell: Acrylic

      - Optional pH Sensor Flow Cell: PVC

    - The sensor interface to the controller is through a digital gateway.

    - The chlorine sensor automatically compensates for temperature utilizing an embedded temperature sensor.

    - The electrolyte provides internal, buffered pH compensation in the range of 4-9 pH units.

    - The sensor includes proprietary Cal Watch self-diagnostic technology.

    - The panel assembly includes a flow cell that integrates a flow meter and control valve.

  - Components

    - Standard equipment:

      - Stainless Steel Mounting Panel

      - Chlorine Sensor with Membrane and Electrolyte

      - Chlorine Sensor flow cell

      - Flow meter with control valve

      - Digital gateway to sc controller with cable

      - User Manual

    - Dimensions

      - Sensor

        - Length: 7.68 in. (195 mm)

        - Diameter: 0.98 in. (25 mm)

      - Panel

        - Length: 19.0 in. (482.6 mm)

        - Width: 19.5 in. (495.3 mm)

        - Depth: 5.95 in. (151.2 mm)

      - Gateway to Controller cable: 3 ft. (1 m)

    - Weight

      - Panel and empty panel-mounted components: approximately 12 lbs. (5.5 kg)

      - Complete panel with pH sensor: approximately 20 lbs. (9.1 kg)

  - Accessories

    - Required

      - sc Controller

    - Optional

      - pH sensor

      - Acidification unit

      - Sample conditioning kit

EXECUTION

  - Preparation

    - Clearances

      - The pre-assembled analyzer panel must be mounted to allow clearance for sensor removal and routine maintenance.

    - Mounting

      - Wall or panel mounted

    - Sample Inlet (order with English or Metric Fittings)

      - English Fittings

        - Speed-fit fitting – ¼ in OD Tube

      - Metric Fittings

        - Speed-fit fitting – 6 mm OD Tube

    - Sample Outlet (order with English or Metric Fittings)

      - English Fittings

        - Speed-fit fitting – ½ in OD

      - Metric Fittings

        - Speed-fit fitting – 12 mm OD

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

- END OF SECTION