PART 1	GENERAL

  - Section includes

    - System that continuously measures conductivity and/or resistivity in aqueous solutions.

  - Measurement Procedures

1.	The method of measurement will be electrodeless/inductive conductivity and resistivity with a probe or sensor.

  - Alternates

    - Probes or sensors that do not communicate with Hach model sc100 or sc1000 controller will not be acceptable.

  - System Description

    - Performance Requirements

      - Measurement range:

        - Conductivity: 0 to 200 microSiemens/cm, up to 0 to 2,000,000 microSiemens/cm

        - Temperature: -10 to 200 °C (14 to 392 °F)

  - Certifications

    - General Purpose CSA/CSANRTL and FM (UL Pending) when used with an approved controller.

    - Class 1, Div 2 Groups A thru D CSA/CSANTRL and FM (UL Pending) when used with an approved controller.

  - Environmental Requirements

    - Operational Criteria

      - Operating temperature: 14 to 392 °F, limited only by sensor body material and mounting hardware

      - Pressure and temperature limits, depending on materials:

        - Polypropylene: 100 psi at 212 °F (6.9 bar at 100 °C)

        - PVDF: 100 psi at 248 °F (6.9 bar at 120 °C)

        - PEEK: 200 psi at 392 °F (13.8 bar at 200 °C)

        - PFA Teflon: 200 psi at 392° F (13.8 bar at 200 °C)

      - Flow rate: 10 ft./s (3 m/s), maximum

  - Warranty

    - The system is warranted for 1 year from date of shipment against defects in materials or workmanship.

  - Maintenance Service

    - Scheduled maintenance:

      - Clean to maintain measurement accuracy. Schedule (days, weeks, etc.) is affected by characteristics of the process solution and should be determined by operating experience.

PART 2	PRODUCTS

  - Manufacturer

    - Hach Company, Loveland, CO

      - Hach 3700-SC Electrodeless Conductivity System

  - Manufactured Unit

    - The Hach 3700-SC Electrodeless Conductivity System consists of:

      - Submersible probe

      - Integral cable

      - Digital gateway device

  - Equipment

    - The 3700-SC system works with Hach models sc100 or sc1000 controllers only.

    - The probe communicates with the controller digitally via RS-485 MODBUS® connection.

    - The probe has a built in Pt 1000 RTD temperature compensator.

    - The probe is water resistant.

    - The wetted material is [select one]:

      - Polypropylene

      - PVDF

      - PEEK®

      - PFA Teflon®

    - Mounting styles [select one]:

# Convertible style:

        - To directly fasten onto the end of a pipe for immersion mounting

        - To mount into any 2-inch NPT fitting with a GLI union-mount adapter

        - To insertion mount into a 2-inch ball valve assembly

- Sanitary style:

  - To conform to provisions of 3-A Sanitary Standards for CIP cleaning

  - With integral 2-inch sanitary mount flange to mount into a standard 2-inch sanitary tee

  - Components

    - Standard equipment:

      - Probe

      - Integral cable

      - Digital gateway

      - Manual

      - Digital gateway

    - Dimensions: dependent on probe selected

    - Weight: dependent on probe selected

  - Accessories

    - Plug in extension cables to extend the distance between the sensor and cable up to 1000 meters (3240 ft.)

    - Junction box for extension cables. Must be used for lengths greater than 100 meters.

    - Extension cables

    - Mounting hardware

PART 3	EXECUTION

  - Preparation

    - The system must be mounted to a Hach mounting assembly directly in the solution to be measured.

  - Installation

    - Contractor will install the analyzer in strict accordance with the manufacturer’s instructions and recommendation.

    - Manufacturer’s representative will include a half-day of start-up service by a factory-trained technician, if requested.

      - Contractor will schedule a date and time for start-up.

      - Contractor will require the following people to be present during the start-up procedure.

        - General contractor

        - Electrical contractor

        - Hach Company factory trained representative

        - Owner’s personnel

        - Engineer

  - Manufacturer’s Service and Start-Up

    - Contractor will include the manufacturer’s services to perform start-up on instrument to include basic operational training and certification of performance of the instrument.

    - Contractor will include a manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment startup through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.

    - Items A and B are to be performed by manufacturer’s factory-trained service personnel.  Field service and factory repair by personnel not employed by the manufacturer is not allowed.

    - Use of manufacturer’s service parts and reagents is required.  Third-party parts and reagents are not approved for use.

END OF SECTION